## Histology Slide ZPL
Testing | What was Done to Test | Result | Fixed if there is a Failure
Testing the edges of the truncation: | set the last name and test name to be right on the edge of the trucation (10-12, 14-16) | 10. success. 11. success. 12. success. 14. success. 15. success.16. success. | N/A
Test to see what happens if a part of the label is left empty: | Set multiple fields as empty | no problems | N/A
Test to see if truncation works when field is set to exactly 13 (the cut off limit) | Set last name and test name to exactly 13 characters | no problems | N/A

## Embedding Dialog
Breast Case List |  |  |  |  | Hold List |  |  |  |  | Not Scanned List
What is being Tested | how it is being tested | Result of Testing | Fixed if Failure Occurs | Date Performed | What is being Tested | how it is being tested | Result of Testing | Fixed if Failure Occurs | Date Performed | What is being Tested | how it is being tested | Result of Testing | Fixed if Failure Occurs | Date Performed
Verify that fixation time is setting correctly | Step 1: searched for breast cases from the last month and went into the report orders specimen details page to check and make sure the fixation information makes sense and matches what it should. Step 2. Check the most recent case on the Embedding log for errors | Error noted…A block from an older case was reprocessed and scanned out at the end of processing…this reset the fixation time for the specimen sa if the entire specimen had been reprocessed and not just the one block. | Since the embedding window is relative now the fixation end date and time will be set back to the day the original specimen was processed to set the specimen fixation back to the correct time. Also a memeting between histo and Sid to determine future resilutions to similar problems is set | 2016-10-17 | Make sure that the blocks/case are being put on hold right when hold is chosen. | Make a MM case and put a block/specimen on hold | worked fine | N/A | 2016-10-17 | check to make sure that cases are added to this list when the case is created. | create a MM case and look to see if the case was added to the not scanned list. | error. Problem with accessioning where the needs gross was not checked | set the check to needs gross and then look into why the needs gross box wasn't checked | 2016-10-17
Verify that collection time is correct | Check the specimen details page and compare the collection time and fixation start time to that noted on the requisition. | No discrepencies noted | N/A | 2016-10-17 | Place on hold from the specimen details page. | try to put a case on hold from the details page. | did not work | Sid notified to be tested again later. | 2016-10-17 | Why was the "needs gross" box not checked on test specimen/patient | create a MM case to see if the needs gross box will stay checked or will be unchecked. | Seems that unless a template is chosen (at least for the Yellowstone Pathologists PC client) the needs gross box is not checked. | need discussion with Sid to understand hwy this is occuring and to come up with a fix. | 2016-10-17
verify that fixation end time is set correctly | Double check the specimen details page and the breast cases list on the embedding dialog to verify that the fixation end time is being correctly set per the informatio ngiven. | When dealing with client accession cases the end time/ fixation duration isn't noted on the Breast Case list in the embedding dialog. | Sid to deal with. | 10/17/2016 … corrected/ annotated following discussion with Sid 10/18/2016 | REPEAT: Try to place a specimen on hold from the details page. | using a MM case go into the specimen details page and using the hold button located there in, place the specimen on hold | worked fine | N/A | 2016-10-17 | Check how the not scanned list is being populated. | create a MM case that is client accesioned, not client accessioned, and has no template to see if there is a point where a case does NOT show up when it should | Both cases showed up in the Not Scanned list | N/A | 2016-10-17
verify that fixation duration is being chosen appropriately | Check the breast cases list for cases with abdnormal fixation duration and check to specimen details page to make sure that the correct proccessor with duration is being chosen correctly. Looking at on the specimens scanned Monday October 17 2016. | Error: as noted in the first stage of testing an error was noted where a reprocessed block was scanned after reprocessing thereby incorrectly setting the fixation duration on the specimen. | processor end time was set back to the original end time for the specimen and the block rescanned. Also a meeting is set for Histo to meet with Sid and discuss possible ways to account for things like added blocks and reprocessed blocks. | 2016-10-17 | try to put a specimen on hold by scanning the container. | scan the container label for a MM case while in the embedding hold list area to place the case on hold. | worked fine | N/A | 2016-10-17 | Checked the Not Scanned List. | Noted that some client accessioned cases were showing up on the list. | These cases show up because the block was not checekd as client accessioned at accessioning. | It will be communicated to accessioning staff that all client accessioned cases must hae the blocks marked as client accessioned to keep these cases from the Not Scanned List. | 2016-10-18
Check to make sure that when a breast case is accessioned/created it is added to the list right away. | A MM breast case will be created to make sure the case is added to the breast case list on emedding dialog | Case was added to the Breast case list as soon as accessing was finished. | N/A | 2016-10-17 | Test to see if putting one block on hold works | A MM case is created with mulitiple blocks, taken to the gross station and one block is put on hold. Press and hold the block you want and select the hold block option | worked fine | N/A | 2016-10-18 | test embedding dialog setting fixation time. | scan a block from yesterday and see if it sets the fixation time with an additional 24 hrs. | worked fine | N/A | 2016-10-18
Check to verify that the fixation duration being listed is also the duration that was entered via the mapping page. | Look at the fixation duration in the embedding diolog and compare the number there to the number entered in the mapping for client accessioned cases. | there is a rounding error that ends with the duration being an hour off one way or the other. There are "0" durations. | Sid to deal with. | 2016-10-18 | Test to see if putting an entire case on hold works | A MM case is created with mulitiple blocks, taken to the gross station and the whole case is put on hold | worked fine | N/A | 2016-10-18 | test embedding dialog setting fixation time. | scan a block from yesterday that has been incorrectly set and properly set the fixation time from the embedding dialog | worked fine | N/A | 2016-10-18
[10 empty] | test embedding dialog setting fixation time. | scan a block from several days ago and see if the fixation time is reset to an incorrect value. | worked fine | N/A | 2016-10-18
[10 empty] | test embedding dialog setting fixation time. | Scan a MM case created today to make sure that the fixation duration is still setting correctly for a specimen that has not been previously scanned. | while setting fixation time worked fine there was the error that allows for negative fixation duration. | Sid to fix. | 2016-10-18
[10 empty] | test embedding dialog setting fixation time. | Scan a case from two days ago to see if fixation duration is updated.